## Summary
Audit Trail Review — Evidence Package
⚠ CRITICAL FINDINGS REQUIRING IMMEDIATE ACTION (1 event)
User | Rule Triggered | Record / Table | Date & Time
analyst_morris | Rule 2 — Audit Trail Integrity Event | CFG-001 | 2024-06-15 09:00:00
REVIEW INFORMATION
System Name | BioLIMSv1
Review Period | 15-Jun-2024 → 15-Jun-2024
Source File | AT_rule02_audit_trail_tampering.csv
Analysis Date (UTC) | 2026-05-06 15:37 UTC
Regulatory Basis | 21 CFR Part 11 §11.10(e) ¦ EU Annex 11, Clause 9 ¦ GAMP 5 (2nd Ed, 2022) ¦ FDA CSA Final Guidance (Sep 2021)
DATASET INTEGRITY
Total Events in Upload | 2
Critical Events | 1
High Events | 0
Medium Events | 0
Auto-Cleared (Low / No Risk) | 1
Events for Review | 1
DATASET HASH (TAMPER EVIDENCE)
SHA-256 (full dataset) | 30703c3c881af2a029009718d60436d42d47a8e72d364d6976564513032e7a97
MD5 (full dataset) | f42e87eb3c75c86387ecc006f71515b1
Hash Scope | All scored columns, sorted alphabetically, CSV-serialized before hashing
REVIEW NARRATIVE
The audit trail for nhn was reviewed from 15-Jun-2024 to 15-Jun-2024. Of 2 recorded events, 1 was escalated for review (1 Critical). Full dataset tier distribution is available in the Full Audit Log sheet. Overall, the system presents a moderate-to-high risk to data integrity. The critical finding identified requires prompt investigation and formal non-conformance documentation.
REVIEWER STATEMENT
This audit trail review was performed using rule-based anomaly detection with system-generated narrative summaries. All risk classifications are derived from deterministic rules; narrative text does not influence risk scoring or tier assignment. 2 records were reviewed across the period 15-Jun-2024 to 15-Jun-2024. 50.0% of records were automatically cleared as low risk. The 1 records requiring review are documented in the 'Events for Review' sheet. Each finding has been independently reviewed and dispositioned by the undersigned reviewer. The 'System-Proposed Disposition' column is informational only; the reviewer's determination in the 'Reviewer Decision' column reflects independent human judgment and is the authoritative record. Regulatory basis: 21 CFR Part 11 §11.10(e) and EU Annex 11 Clause 9. Timezone note: Off-hours scoring (Rule 10) assumes all timestamps in the uploaded file are in the same timezone. If your system exports timestamps in UTC or a non-local timezone, off-hours flags should be interpreted accordingly. Reviewers should verify local shift patterns before dispositing temporal anomalies.
Reviewer Name
Reviewer Title
Date of Review
Reviewer Signature

## Events for Review
No. | Events (rolled up) | Risk Level | Date & Time | User | Action | Record | Record ID | Change Reason | Primary Rule | Related Sequence | Why It Matters | What Happened | Recommended Action | System-Proposed Disposition (see Decision Basis — reviewer must make independent determination) | Decision Basis | Reviewer Decision (independent — tick one) | Reviewer Notes
1 | 1 | Critical | 2024-06-15 09:00:00 | analyst_morris | UPDATE | AUDIT_CONFIG | CFG-001 | Adjusted retention setting per IT request | Rule 2 — Audit Trail Integrity Event [CRITICAL] |  | 21 CFR Part 11 §11.10(e); EU Annex 11 Clause 9 — Audit trail systems must be protected from modification; any change to audit trail configuration is a critical integrity event. | analyst_morris performed UPDATE on AUDIT_CONFIG/CFG-001 at 2024-06-15 09:00:00; comment on file reads 'Adjusted retention setting per IT request' — a direct write to the audit trail configuration or records is a critical integrity event that must be formally explained and authorized. | This action affected the audit trail system itself. Retrieve the full system configuration log and determine the scope of the change. Assess whether any events may have been suppressed during the affected period. Document all findings and escalate immediately. |  |  | ☐ Justified ☐ Escalate to CAPA ☐ False Positive
Note on A1: VALINTEL.AI: Showing 1 escalated event(s) from the Full Audit Log (1 Critical). Each event triggered at least one named detection rule (Rules 1–15) at or above its threshold. All events remain visible in the Full Audit Log sheet.

## Full Audit Log
Date & Time | User | Action | Record Type | Record ID | Old Value | New Value | Change Reason | Risk Level | Primary Rule | All Rules Fired | Related Sequence
2024-06-15 09:00:00 | analyst_morris | UPDATE | AUDIT_CONFIG | CFG-001 | retention_days=365 | retention_days=180 | Adjusted retention setting per IT request | Critical | Rule 2 — Audit Trail Integrity Event [CRITICAL] | Rule 2 — Audit Trail Integrity Event
2024-06-15 10:00:00 | analyst_nguyen | UPDATE | RESULTS | RES-9001 | 6.8 | 7.0 | pH recalibrated per SOP-QC-022 after instrument service | Low | No named rule threshold met | No anomaly detected

## Detection Logic
# | Tier | Class | Rule Name | Trigger Condition | FDA Regulation | EU Annex 11 | Status
1 | Critical | Tier 1 | Record Reconstruction | Same user deletes then recreates the same record_id within 4 hours. | 21 CFR Part 11 §11.10(e); ALCOA+ Original | Clause 9 | ACTIVE
2 | Critical | Tier 1 | Audit Trail Tampering | Any action on audit trail config table. High: DELETE on GxP-sensitive record type. | 21 CFR Part 11 §11.10(e) | Clause 9 | ACTIVE
3 | High | Tier 1 | Missing Before/After Values | UPDATE: old_value = new_value or either null. CREATE: new_value null. DELETE: old_value null. Skipped if cols absent. | 21 CFR Part 11 §11.10(e); ALCOA+ Original | Clause 9 | ACTIVE
4 | Critical | Tier 1 | Self-Approval SoD Violation | Same user appears as creator AND approver on the same record_id. | 21 CFR Part 11 §11.10(d); EU Annex 11 Clause 12 | Clause 12 | ACTIVE
5 | Critical | Tier 1 | Modification After Approval | UPDATE or DELETE on a record_id after an APPROVE or RELEASE action already exists for that record. | 21 CFR Part 11 §11.10(e); ALCOA+ Original | Clause 9 | ACTIVE
6 | High | Tier 1 | Vague Rationale | Fires on UPDATE/DELETE where the change reason (comments column) is blank OR a single generic word from the blocklist. EXEMPT when comment has 3+ words AND contains at least one domain term. BLOCKLIST: error, fix, fixed, corrected, done, ok, okay, updated, changed, modified, n/a, na, per sop (without specifics), human error, training issue, typo, mistake, correction, redo. DOMAIN TERMS (auto-exempt): pH, temperature, calibration, batch, specification, deviation, OOS, out-of-spec, protocol, SOP-[ref], IQ, OQ, PQ, CAPA, method, stability, sterility, endotoxin, assay, analyte, sample, reagent, standard, yield, potency, identity, purity. EXAMPLES NOT firing: 'pH corrected again' (pH = domain term), 'data correction on repeat experiment' (5 words + correction+experiment), 'per SOP-QC-014 §4.2' (SOP reference). EXAMPLES THAT FIRE: blank, 'Fixed', 'Error', 'Done', 'OK', 'Per SOP' (no number), 'Human error', 'Training issue'. | 21 CFR Part 211.68; ALCOA+ Attributable and Legible | Clause 9 | ACTIVE
7 | High | Tier 1 | Statistical Value Outlier | Modification to GxP-critical field AND change_reason blank AND approved_by blank (all three conditions). Does NOT fire if either documentation field is populated. | 21 CFR 211.68(b); EU Annex 11 §10; ALCOA+ Attributable + Original | Clause 10 | ACTIVE
8 | Critical | Tier 1 | Service / Shared Account GxP Action | Non-personal account (svc_, batch_, api_, root, daemon, guest, test_) performs GxP data action. | 21 CFR Part 11 §11.300 | Clause 12 | ACTIVE
9 | High | Tier 1 | Failed Login → Data Manipulation (narrowed) | AUTH_FAILURE event followed by write action by same user_id within 15 minutes on a GxP-critical table. Silent-skip if dataset has zero auth events. | 21 CFR 11.10(d); 21 CFR 211.68; ALCOA+ Attributable | Clause 12 | ACTIVE
10 | High | Tier 1 | Privileged User GxP Action | Role contains admin/system/superuser/root/sa/sysadmin/dba AND action_type ∈ {INSERT,UPDATE,DELETE,MODIFY} AND record_type is GxP-critical. Single clean check — old Rule 3/8 distinction removed. | 21 CFR 11.10(c)/(d); EU Annex 11 §2; ICH Q9 | Clause 12 | ACTIVE
11 | High | Tier 1 | Missing User Attribution | Null or blank user_id on any audit trail event. | 21 CFR Part 11 §11.10(e); ALCOA+ Attributable | — | ACTIVE
12 | Critical | Tier 1 | Timestamp Reversal | Approval timestamp earlier than creation timestamp for same record_id. Chronologically impossible. | 21 CFR Part 11 §11.10(e); ALCOA+ Contemporaneous | Clause 9 | ACTIVE
13 | High | Tier 1 | Missing Timestamp | Null or unparseable timestamp on any audit trail event. | 21 CFR Part 11 §11.10(e); ALCOA+ Contemporaneous | — | ACTIVE
15 | Medium | Tier 1 | Contemporaneous Burst (narrowed) | ≥8 events by same human user_id within a 10-minute rolling window on GxP-critical records only. Service/shared accounts excluded. | ALCOA+ Contemporaneous; 21 CFR Part 11 §11.10(e); EU Annex 11 §9 | Clause 9 | ACTIVE
17 | Critical | Tier 1 | Workflow Status Reversal | Backward workflow status transition on same record_id (e.g. Approved→Draft, Released→Pending, Closed→Open). Fires Critical when record is GxP-critical (RESULTS, BATCH, etc.); High otherwise. Silent-skip if no status/state/workflow column present. | 21 CFR Part 11 §11.10(e); ALCOA+ Original; GAMP 5 (2nd Ed, 2022) | Clause 9 | ACTIVE

## Rule Summary
Rule Firing Summary — Events Escalated per Detection Rule
Based on 2 total events · 1 escalated · System: nhn
# | Rule Name | Tier | Count | % Share
1 | Record Reconstruction | Critical | 0 | 0.0%
2 | Audit Trail Tampering | Critical | 1 | 100.0%
3 | Missing Before/After Values | High | 0 | 0.0%
4 | Self-Approval SoD Violation | Critical | 0 | 0.0%
5 | Modification After Approval | Critical | 0 | 0.0%
6 | Vague Rationale | High | 0 | 0.0%
7 | Statistical Value Outlier | High | 0 | 0.0%
8 | Service / Shared Account GxP Action | Critical | 0 | 0.0%
9 | Failed Login → Data Manipulation | High | 0 | 0.0%
10 | Privileged User GxP Action | High | 0 | 0.0%
11 | Missing User Attribution | High | 0 | 0.0%
12 | Timestamp Reversal | Critical | 0 | 0.0%
13 | Missing Timestamp | High | 0 | 0.0%
15 | Contemporaneous Burst | Medium | 0 | 0.0%
17 | Workflow Status Reversal | Critical | 0 | 0.0%
Count = events where this rule is the Primary Rule. One event may trigger multiple rules — only the highest-priority rule is counted here. Low-risk (auto-cleared) events are excluded.